## Sequences
ID codes for States (ST-) and Sequences (SEQ-) |  |  |  |  | Sequences: defined order for operating different states
ID | Full ID | Type | ID no. | State title | ID | Description | Instruction | Duration | Notes
ST-01 | ST-01-LEM reservoir fill | ST | 01 | LEM reservoir fill | SEQ-01-Drain the system |  |  |  |  | Build out in terms of desired set point (alpha), then later (not alpha) alarm high/ high high and low/ low low, action e.g. notify/Stop
ST-02.1 | ST-02.1-Line fill L04 and L11 | ST | 02.1 | Line fill L04 and L11 | ST-03.5 | Drain ICS | Complete then proceed | Minutes
ST-02.2 | ST-02.2-Line fill ECS | ST | 02.2 | Line fill ECS | ST-03.2 | Drain step 2 | Complete then proceed | Minutes
ST-02.3 | ST-02.3-Line fill ICS | ST | 02.3 | Line fill ICS | ST-03.3 | Drain step 3 | Complete then proceed | Minutes
ST-02.4 | ST-02.4-Line fill and circulation | ST | 02.4 | Line fill and circulation | ST-03.4 | Drain to empty the reservoir | Complete | Minutes |  | Parameters
ST-03.1 | ST-03.1-Drain circulation loop | ST | 03.1 | Drain circulation loop | SEQ-02-Sterilisation
ST-03.2 | ST-03.2-Drain step 2 | ST | 03.2 | Drain step 2 | SEQ-01 | Drain the system | Complete then proceed | Minutes |  | "check it is complete"
ST-03.3 | ST-03.3-Drain step 3 | ST | 03.3 | Drain step 3 | ST-01 | LEM reservoir fill | Complete then proceed | Minutes | Fill with sterilant | Pressure
ST-03.4 | ST-03.4-Drain to empty the reservoir | ST | 03.4 | Drain to empty the reservoir | ST-02.1 | Line fill L04 and L11 | Complete then proceed | Minutes |  | AT01 level
ST-03.5 | ST-03.5-Drain ICS | ST | 03.5 | Drain ICS | ST-02.2 | Line fill ECS | Complete then proceed | Minutes |  | Time
ST-04 | ST-04-Sampling | ST | 04 | Sampling | ST-02.3 | Line fill ICS | Complete then proceed | Minutes |  | Volume dose??
ST-05 | ST-05-Inoculation | ST | 05 | Inoculation | ST-02.4 | Line fill and circulation | Run for defined time period | Hours
ST-06 | ST-06-Dynamic seeding | ST | 06 | Dynamic seeding | SEQ-03-Wash
ST-07 | ST-07-ECS liquid drain/replacement (remove unattached cells) | ST | 07 | ECS liquid drain/replacement (remove unattached cells) | SEQ-01 | Drain the system | Complete then proceed | Minutes
ST-08.1 | ST-08.1-Liq circ (ICS inlet to ICS + ECS outlet only) | ST | 08.1 | Liq circ (ICS inlet to ICS + ECS outlet only) | ST-01 | LEM reservoir fill | Complete then proceed | Minutes | Fill with wash liquid
ST-08.2 | ST-08.2-Liq circ (ICS inlet to ICS outlet only) | ST | 08.2 | Liq circ (ICS inlet to ICS outlet only) | ST-02.1 | Line fill L04 and L11 | Complete then proceed | Minutes
ST-08.3 | ST-08.3-Liq circ (ICS inlet to ECS outlet only) | ST | 08.3 | Liq circ (ICS inlet to ECS outlet only) | ST-02.2 | Line fill ECS | Complete then proceed | Minutes
ST-09.1a | ST-09.1a-Liq circ (ICS + ECS inlet to ICS + ECS outlet) - option 1 | ST | 09.1a | Liq circ (ICS + ECS inlet to ICS + ECS outlet) - option 1 | ST-02.3 | Line fill ICS | Complete then proceed | Minutes
ST-09.1b | ST-09.1b-Liq circ (ICS + ECS inlet to ICS + ECS outlet) - option 2 | ST | 09.1b | Liq circ (ICS + ECS inlet to ICS + ECS outlet) - option 2 | ST-02.4 | Line fill and circulation | Run for defined time period | Hours
ST-09.2 | ST-09.2-Liq circ (ICS + ECS inlet to ICS outlet only) | ST | 09.2 | Liq circ (ICS + ECS inlet to ICS outlet only) | SEQ-04-Pretreat
ST-09.3 | ST-09.3-Liq circ (ICS + ECS inlet to ECS outlet only) | ST | 09.3 | Liq circ (ICS + ECS inlet to ECS outlet only) | SEQ-01 | Drain the system | Complete then proceed | Minutes
ST-10.1a | ST-10.1a-Liq circ (ECS inlet to ICS + ECS outlet only) | ST | 10.1a | Liq circ (ECS inlet to ICS + ECS outlet only) | ST-01 | LEM reservoir fill | Complete then proceed | Minutes | Fill with pretreat liquid
ST-10.1b | ST-10.1b-Liq circ (ECS inlet to ICS + ECS outlet only) | ST | 10.1b | Liq circ (ECS inlet to ICS + ECS outlet only) | ST-02.4 | Line fill and circulation | Run for defined time period | Hours | Assume only need to fill the main circulation loop with pretreatment liquid
ST-10.2 | ST-10.2-Liq circ (ECS inlet to ICS outlet only) | ST | 10.2 | Liq circ (ECS inlet to ICS outlet only) | SEQ-05-Prime the system prior to inoculuation
ST-10.3 | ST-10.3-Liq circ (ECS inlet to ECS outlet only) | ST | 10.3 | Liq circ (ECS inlet to ECS outlet only) | SEQ-01 | Drain the system | Complete then proceed | Minutes
ST-11.1 | ST-11.1-Harvest loop fill | ST | 11.1 | Harvest loop fill | ST-01 | LEM reservoir fill | Complete then proceed | Minutes | Fill with fresh culture medium
ST-11.2 | ST-11.2-Harvest loop in situ circulation | ST | 11.2 | Harvest loop in situ circulation | ST-02.4 | Line fill and circulation | Complete | Minutes | Assume only need to fill the main circulation loop with desired liquid
ST-11.3 | ST-11.3-Harvest loop drain & cell removal | ST | 11.3 | Harvest loop drain & cell removal | SEQ-06-Seeding
ST-11.4 | ST-11.4-Harvest loop flush | ST | 11.4 | Harvest loop flush | SEQ-05 | Prime the system prior to inoculuation | Complete then proceed | Minutes
SEQ-01 | SEQ-1-Drain the system | SEQ | 1 | Drain the system | ST-05 | Inoculation | Complete then proceed | Minutes
SEQ-02 | SEQ-2-Sterilisation | SEQ | 2 | Sterilisation | ST-06 | Dynamic seeding | Run for defined time period | Hours | Isolation of reactor with defined motion sequence
SEQ-03 | SEQ-3-Wash | SEQ | 3 | Wash | ST-07 | ECS liquid drain/replacement (remove unattached cells) | Complete | Minutes | Replace the ECS with fresh culture media and remove unattached cells
SEQ-04 | SEQ-4-Pretreat | SEQ | 4 | Pretreat | SEQ-07-Circulation and culture period
SEQ-05 | SEQ-5-Prime the system prior to inoculuation | SEQ | 5 | Prime the system prior to inoculuation | SEQ-06 | Seeding | Check it has been completed
SEQ-06 | SEQ-6-Seeding | SEQ | 6 | Seeding | ST-08.1 | Liq circ (ICS inlet to ICS + ECS outlet only) | Run for defined time period | Days | Need to initiate control loops and control do set point parameters
SEQ-07 | SEQ-7-Circulation and culture period | SEQ | 7 | Circulation and culture period | SEQ-08-Sampling
SEQ-08 | SEQ-8-Sampling | SEQ | 8 | Sampling | SEQ-07 | Circulation and culture period | Check it has been initiated
SEQ-09 | SEQ-9-Media addition during culture | SEQ | 9 | Media addition during culture | ST-08.1 | Liq circ (ICS inlet to ICS + ECS outlet only) | Continue running in parallel
SEQ-10 | SEQ-10-Media exchange (partial, HFB remains full) | SEQ | 10 | Media exchange (partial, HFB remains full) | ST-04 | Sampling | Complete | Minutes
SEQ-11 | SEQ-11-Media exchange (full, HFB drained of liquid) | SEQ | 11 | Media exchange (full, HFB drained of liquid) | SEQ-09-Media addition during culture
SEQ-12 | SEQ-12-In situ harvest | SEQ | 12 | In situ harvest | SEQ-07 | Circulation and culture period | Check it has been initiated
[5 empty] | ST-08.1 | Liq circ (ICS inlet to ICS + ECS outlet only) | Continue running in parallel
[5 empty] | ST-01 | LEM reservoir fill | Complete | Minutes
SEQ-10-Media exchange (partial, HFB remains full)
[5 empty] | SEQ-07 | Circulation and culture period | Check it has been initiated
[5 empty] | ST-08.1 | Liq circ (ICS inlet to ICS + ECS outlet only) | Pause
[5 empty] | ST-03.4 | Drain to empty the reservoir | Complete then proceed | Minutes
[5 empty] | ST-03.3 | Drain step 3 | Complete then proceed | Minutes
[5 empty] | ST-01 | LEM reservoir fill | Complete then proceed | Minutes
[5 empty] | ST-08.1 | Liq circ (ICS inlet to ICS + ECS outlet only) | Restart | Days
SEQ-11-Media exchange (full, HFB drained of liquid)
[5 empty] | SEQ-07 | Circulation and culture period | Check it has been initiated
[5 empty] | ST-08.1 | Liq circ (ICS inlet to ICS + ECS outlet only) | Pause
[5 empty] | ST-03.1 | Drain circulation loop | Complete then proceed | Minutes
[5 empty] | ST-03.2 | Drain step 2 | Complete then proceed | Minutes
[5 empty] | ST-03.3 | Drain step 3 | Complete then proceed | Minutes
[5 empty] | ST-03.4 | Drain to empty the reservoir | Complete then proceed | Minutes
[5 empty] | ST-01 | LEM reservoir fill | Complete then proceed | Minutes
[5 empty] | ST-07 | ECS liquid drain/replacement (remove unattached cells) | Complete then proceed | Minutes
[5 empty] | ST-08.1 | Liq circ (ICS inlet to ICS + ECS outlet only) | Restart | Days
SEQ-12-In situ harvest
[5 empty] | SEQ-01 | Drain the system | Complete then proceed | Minutes
[5 empty] | ST-01 | LEM reservoir fill | Complete then proceed | Minutes | Fill with wash agent
[5 empty] | ST-02.4 | Line fill and circulation | Complete then proceed | Minutes
[5 empty] | SEQ-01 | Drain the system | Complete then proceed | Minutes
[5 empty] | ST-01 | LEM reservoir fill | Complete then proceed | Minutes | Fill with dissociation agent
[5 empty] | ST-11.1 | Harvest loop fill | Complete then proceed | Minutes
[5 empty] | ST-11.2 | Harvest loop in situ circulation | Complete then proceed | Minutes
[5 empty] | ST-11.3 | Harvest loop drain & cell removal | Complete then proceed | Minutes
[5 empty] | ST-01 | LEM reservoir fill | Complete then proceed | Minutes | Fill with wash agent
[5 empty] | ST-11.4 | Harvest loop flush | Complete then proceed | Minutes
[5 empty] | SEQ-12 | #N/A | Repeat if required
Note on A18: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Grey text = non-standard operating mode.

## Valve and pump states
States define the required "state change". i.e. if a state does not denote a valve as open or close then the previous state remains active.
State |  |  | P01 | P02 | V01 | V02 | V03 | V04 | V05 | V06 | V07 | V08 | V09 | V10 | V11 | V12 | [5 empty] | ID | P01 | P02 | V01 | V02 | V03 | V04 | V05 | V06 | V07 | V08 | V09 | V10 | V11 | V12 | [5 empty] | P1 (ml/min) | P2 (ml/min) | [7 empty] | Actually NC - 0 & 1 swapped here |  | Actually W - does it need to be swapped? |  | Actually W - does it need to be swapped? | [6 empty] | Not used | Not used | Not used | Not used | [10 empty] | Transition - Time |  | Transition - Temp |  | Transition - DO
|  | Natural state (NO or NC) |  |  | NO | NO | NO | NC | NO | NC | NO | NC | NC | NC | NO | NC | [8 empty] | Rear | Front | Rear | Front | Rear | Front | Rear | Front | Front | Front | [15 empty] | NO | NC | W | W | NO | W | W | W | W | NC
|  | Actual natural state (NO or NC) |  |  | NO | NC | W | W | NO | W | W | W | W | NC | [10 empty] | NO | NC | W | W | NO | W | W | W | W | NC |  |  |  | NULLLEADER | MODID | COMMAND | STATEID | OXYCIRCFLAG | PRESSUREPUMPFLAG | DISPENSEPARA | DISPENSEVOLUMESP | NA | NA | NA | FLAGCHECK | VALVE1 | VALVE2 | VALVE3 | VALVE4 | VALVE5 | VALVE6 | VALVE7 | VALVE8 | VALVE9 | VALVE10 | VALVE11 | VALVE12 | VALVE13 | VALVE14 | VALVE15 | VALVE16 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | TRANSTIMEFLAG | TRANSTIMESP | TRANSTEMPFLAG | TRANSTEMPSP | TRANSOXYFLAG | TRANSOXYSP | SERIALERRORFLAG | TRANSTYPE | FLOWINSTALLED | PRESSUREINSTALLED | OXYINSTALLED | CIRCPUMPCAL | PRESSUREPUMPCAL | NA | PCBAUDRATE | OXYBAUDRATE | PRESSUREBAUDRATE | VALVE1PIN | VALVE2PIN | VALVE3PIN | VALVE4PIN | VALVE5PIN | VALVE6PIN | VALVE7PIN | VALVE8PIN | VALVE9PIN | VALVE10PIN | VALVE11PIN | VALVE12PIN | VALVE13PIN | VALVE14PIN | VALVE15PIN | VALVE16PIN | PUMP1DIRPIN | PUMP1STEPPIN | PUMP2DIRPIN | PUMP2STEPPIN | PUMP3DIRPIN | PUMP3STEPPIN | PUMP4DIRPIN | PUMP4STEPPIN | OXYCONFIGPARAMETER | PRESSURE0ADDRESS | OXYCIRCCHANNEL | PRESSURESP | PRESSUREKP | PRESSUREKI | PRESSUREKD | PRESSURECHANNEL | ACTIVEOXYCHANNEL1 | ACTIVEOXYCHANNEL2 | ACTIVEOXYCHANNEL3 | ACTIVEOXYCHANNEL4 | FLOW0ADDRESS | OXYSP | OXYKP | OXYKI | OXYKD | NA | NA | NULLTRAILER
ST-01 |  | LEM reservoir fill | Off | Off | Natural | Natural | Natural | Natural | Natural | Natural | Natural | Natural | Natural | Natural | Natural | Natural | [5 empty] | 4010 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | LEM reservoir fill | 0 | 2501 | 1002 | 4010 | 5 | 0 | 0 | 0 | 0 | 0 | 0 | 4010 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 2500 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 1 | 0.003125 | 0.003125 | 0 | 0 | 115200 | 0 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 47 | 4 | 1 | 200 | 2 | 5 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 20 | 0.2 | 0 | 0 | 0 | 0 | 0
ST-02.1 |  | Line fill L04 and L11 | On | Off | Closed | Natural | Natural | Open | Closed | Natural | Closed | Open | Natural | Natural | Open | Closed | [5 empty] | 4021 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | Line fill L04 and L11 | 0 | 2501 | 1002 | 4021 | 40 | 0 | 0 | 0 | 0 | 0 | 0 | 4021 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 10000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 1 | 0.003125 | 0.003125 | 0 | 0 | 115200 | 0 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 47 | 4 | 1 | 200 | 2 | 5 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 20 | 0.2 | 0 | 0 | 0 | 0 | 0
ST-02.2 |  | Line fill ECS | On | On | Closed | Closed | Closed | Open | Open | Open | Closed | Closed | Open | Closed | Open | Closed |  | Key | Description | NO | NC | 4022 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | Line fill ECS | 0 | 2501 | 1002 | 4022 | 40 | 40 | 0 | 0 | 0 | 0 | 0 | 4022 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 10000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 1 | 0.003125 | 0.003125 | 0 | 0 | 115200 | 0 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 47 | 4 | 1 | 200 | 2 | 5 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 20 | 0.2 | 0 | 0 | 0 | 0 | 0
ST-02.3 |  | Line fill ICS | On | On | Open | Open | Closed | Closed | Open | Closed | Open | Closed | Closed | Open | Open | Closed |  | Natural | valve unpowered | 0 | 0 | 4023 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | Line fill ICS | 0 | 2501 | 1002 | 4023 | 100 | 50 | 0 | 0 | 0 | 0 | 0 | 4023 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 40000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 1 | 0.003125 | 0.003125 | 0 | 0 | 115200 | 0 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 47 | 4 | 1 | 200 | 2 | 5 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 20 | 0.2 | 0 | 0 | 0 | 0 | 0
ST-02.4 |  | Line fill and circulation | On | On | Open | Open | Open | Closed | Open | Closed | Open | Closed | Closed | Closed | Open | Closed |  | Open | valve is open | 0 | 1 | 4024 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Line fill and circulation | 0 | 2501 | 1002 | 4024 | 100 | 50 | 0 | 0 | 0 | 0 | 0 | 4024 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 40000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 1 | 0.003125 | 0.003125 | 0 | 0 | 115200 | 0 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 47 | 4 | 1 | 200 | 2 | 5 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 20 | 0.2 | 0 | 0 | 0 | 0 | 0
ST-03.1 |  | Drain circulation loop | On | On | Open | Open | Open | Closed | Open | Closed | Open | Closed | Closed | Closed | Closed | Open |  | Closed | valve is closed | 1 | 0 | 4031 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | Drain circulation loop | 0 | 2501 | 1002 | 4031 | 60 | 30 | 0 | 0 | 0 | 0 | 0 | 4031 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 20000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 1 | 0.003125 | 0.003125 | 0 | 0 | 115200 | 0 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 47 | 4 | 1 | 200 | 2 | 5 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 20 | 0.2 | 0 | 0 | 0 | 0 | 0
ST-03.2 |  | Drain step 2 | On | On | Closed | Closed | Closed | Open | Open | Open | Closed | Closed | Open | Closed | Closed | Open |  | Rev | reversed |  |  | 4032 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | Drain step 2 | 0 | 2501 | 1002 | 4032 | 60 | 60 | 0 | 0 | 0 | 0 | 0 | 4032 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 10000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 1 | 0.003125 | 0.003125 | 0 | 0 | 115200 | 0 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 47 | 4 | 1 | 200 | 2 | 5 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 20 | 0.2 | 0 | 0 | 0 | 0 | 0
ST-03.3 |  | Drain step 3 | On | Off | Closed | Natural | Natural | Open | Closed | Natural | Closed | Open | Natural | Closed | Closed | Open | [5 empty] | 4033 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | Drain step 3 | 0 | 2501 | 1002 | 4033 | 60 | 0 | 0 | 0 | 0 | 0 | 0 | 4033 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 10000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 1 | 0.003125 | 0.003125 | 0 | 0 | 115200 | 0 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 47 | 4 | 1 | 200 | 2 | 5 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 20 | 0.2 | 0 | 0 | 0 | 0 | 0
ST-03.4 |  | Drain to empty the reservoir | On | Off | Closed | Natural | Closed | Open | Closed | Natural | Closed | Closed | Open | Open | Open | Closed | [5 empty] | 4034 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | Drain to empty the reservoir | 0 | 2501 | 1002 | 4034 | 60 | 0 | 0 | 0 | 0 | 0 | 0 | 4034 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 20000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 1 | 0.003125 | 0.003125 | 0 | 0 | 115200 | 0 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 47 | 4 | 1 | 200 | 2 | 5 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 20 | 0.2 | 0 | 0 | 0 | 0 | 0
ST-03.5 |  | Drain ICS | On | On | Open | Open | Closed | Closed | Closed | Closed | Natural | Natural | Open | Closed | Open | Closed | [5 empty] | 4035 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Drain ICS | 0 | 2501 | 1002 | 4035 | 100 | 100 | 0 | 0 | 0 | 0 | 0 | 4035 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 180000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 1 | 0.003125 | 0.003125 | 0 | 0 | 115200 | 0 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 47 | 4 | 1 | 200 | 2 | 5 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 20 | 0.2 | 0 | 0 | 0 | 0 | 0
ST-04 |  | Sampling | [19 empty] | 4040 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | Sampling | 0 | 2501 | 1002 | 4040 | 5 | #N/A | 0 | 0 | 0 | 0 | 0 | 4040 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 2500 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 1 | 0.003125 | 0.003125 | 0 | 0 | 115200 | 0 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 47 | 4 | 1 | 200 | 2 | 5 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 20 | 0.2 | 0 | 0 | 0 | 0 | 0
ST-05 |  | Inoculation | Off | Off | Closed | Closed | Natural | Natural | Open | Open | Closed | Open | Natural | Closed | Natural | Natural | [5 empty] | 4050 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | Inoculation | 0 | 2501 | 1002 | 4050 | 5 | 0 | 0 | 0 | 0 | 0 | 0 | 4050 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 2500 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 1 | 0.003125 | 0.003125 | 0 | 0 | 115200 | 0 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 47 | 4 | 1 | 200 | 2 | 5 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 20 | 0.2 | 0 | 0 | 0 | 0 | 0
ST-06 |  | Dynamic seeding | Off | Off | Closed | Closed | Natural | Natural | Closed | Closed | Natural | Natural | Natural | Natural | Natural | Natural | [5 empty] | 4060 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Dynamic seeding | 0 | 2501 | 1002 | 4060 | 5 | 0 | 0 | 0 | 0 | 0 | 0 | 4060 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 2500 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 1 | 0.003125 | 0.003125 | 0 | 0 | 115200 | 0 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 47 | 4 | 1 | 200 | 2 | 5 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 20 | 0.2 | 0 | 0 | 0 | 0 | 0
ST-07 |  | ECS liquid drain/replacement (remove unattached cells) | On | On | Closed | Closed | Closed | Open | Open | Open | Natural | Natural | Open | Closed | Open | Closed | [5 empty] | 4070 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | ECS liquid drain/replacement (remove unattached cells) | 0 | 2501 | 1002 | 4070 | 5 | 5 | 0 | 0 | 0 | 0 | 0 | 4070 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 2500 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 1 | 0.003125 | 0.003125 | 0 | 0 | 115200 | 0 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 47 | 4 | 1 | 200 | 2 | 5 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 20 | 0.2 | 0 | 0 | 0 | 0 | 0
ST-08.1 |  | Liq circ (ICS inlet to ICS + ECS outlet only) | On | On | Open | Open | Open | Closed | Open | Closed | Open | Closed | Closed | Closed | Open | Closed | default state (most frequent) |  |  |  |  | 4081 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Liq circ (ICS inlet to ICS + ECS outlet only) | 0 | 2501 | 1002 | 4081 | 5 | 5 | 0 | 0 | 0 | 0 | 0 | 4081 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 2500 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 1 | 0.003125 | 0.003125 | 0 | 0 | 115200 | 0 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 47 | 4 | 1 | 200 | 2 | 5 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 20 | 0.2 | 0 | 0 | 0 | 0 | 0
ST-08.2 |  | Liq circ (ICS inlet to ICS outlet only) | On | On | Open | Open | Open | Closed | Closed | Closed | Natural | Natural | Closed | Closed | Open | Closed | [5 empty] | 4082 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Liq circ (ICS inlet to ICS outlet only) | 0 | 2501 | 1002 | 4082 | 60 | 60 | 0 | 0 | 0 | 0 | 0 | 4082 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 10000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 1 | 0.003125 | 0.003125 | 0 | 0 | 115200 | 0 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 47 | 4 | 1 | 200 | 2 | 5 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 20 | 0.2 | 0 | 0 | 0 | 0 | 0
ST-08.3 |  | Liq circ (ICS inlet to ECS outlet only) | On | Off | Open | Closed | Natural | Closed | Open | Closed | Open | Closed | Natural | Closed | Open | Closed | [5 empty] | 4083 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Liq circ (ICS inlet to ECS outlet only) | 0 | 2501 | 1002 | 4083 | 5 | 0 | 0 | 0 | 0 | 0 | 0 | 4083 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 2500 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 1 | 0.003125 | 0.003125 | 0 | 0 | 115200 | 0 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 47 | 4 | 1 | 200 | 2 | 5 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 20 | 0.2 | 0 | 0 | 0 | 0 | 0
ST-09.1a | a | Liq circ (ICS + ECS inlet to ICS + ECS outlet) - option 1 | Rev | Rev | Open | Open | Open | Closed | Open | Open | Open | Closed | Closed | Closed | Open | Closed | [5 empty] | 4091a | -1 | -1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | Liq circ (ICS + ECS inlet to ICS + ECS outlet) - option 1 | 0 | 2501 | 1002 | 4091a | 5 | 5 | 0 | 0 | 0 | 0 | 0 | 4091a | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 2500 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 1 | 0.003125 | 0.003125 | 0 | 0 | 115200 | 0 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 47 | 4 | 1 | 200 | 2 | 5 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 20 | 0.2 | 0 | 0 | 0 | 0 | 0
ST-09.1b | b | Liq circ (ICS + ECS inlet to ICS + ECS outlet) - option 2 | On | On | Open | Open | Open | Open | Open | Open | Closed | Closed | Closed | Closed | Open | Closed | [5 empty] | 4091b | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | Liq circ (ICS + ECS inlet to ICS + ECS outlet) - option 2 | 0 | 2501 | 1002 | 4091b | 5 | 5 | 0 | 0 | 0 | 0 | 0 | 4091b | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 2500 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 1 | 0.003125 | 0.003125 | 0 | 0 | 115200 | 0 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 47 | 4 | 1 | 200 | 2 | 5 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 20 | 0.2 | 0 | 0 | 0 | 0 | 0
ST-09.2 |  | Liq circ (ICS + ECS inlet to ICS outlet only) | On | On | Open | Open | Open | Open | Open | Closed | Closed | Closed | Closed | Closed | Open | Closed | [5 empty] | 4092 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | Liq circ (ICS + ECS inlet to ICS outlet only) | 0 | 2501 | 1002 | 4092 | 5 | 5 | 0 | 0 | 0 | 0 | 0 | 4092 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 2500 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 1 | 0.003125 | 0.003125 | 0 | 0 | 115200 | 0 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 47 | 4 | 1 | 200 | 2 | 5 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 20 | 0.2 | 0 | 0 | 0 | 0 | 0
ST-09.3 |  | Liq circ (ICS + ECS inlet to ECS outlet only) | On | On | Open | Closed | Open | Open | Open | Open | Closed | Closed | Closed | Closed | Open | Closed | [5 empty] | 4093 | 1 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | Liq circ (ICS + ECS inlet to ECS outlet only) | 0 | 2501 | 1002 | 4093 | 5 | 5 | 0 | 0 | 0 | 0 | 0 | 4093 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 2500 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 1 | 0.003125 | 0.003125 | 0 | 0 | 115200 | 0 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 47 | 4 | 1 | 200 | 2 | 5 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 20 | 0.2 | 0 | 0 | 0 | 0 | 0
ST-10.1a | a | Liq circ (ECS inlet to ICS + ECS outlet only) | Rev | Rev | Open | Closed | Open | Closed | Open | Open | Open | Closed | Closed | Closed | Open | Closed | [5 empty] | 4101a | -1 | -1 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | Liq circ (ECS inlet to ICS + ECS outlet only) | 0 | 2501 | 1002 | 4101a | 5 | 5 | 0 | 0 | 0 | 0 | 0 | 4101a | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 2500 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 1 | 0.003125 | 0.003125 | 0 | 0 | 115200 | 0 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 47 | 4 | 1 | 200 | 2 | 5 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 20 | 0.2 | 0 | 0 | 0 | 0 | 0
ST-10.1b | b | Liq circ (ECS inlet to ICS + ECS outlet only) | On | On | Closed | Open | Open | Open | Open | Open | Closed | Closed | Closed | Closed | Open | Closed | [5 empty] | 4101b | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | Liq circ (ECS inlet to ICS + ECS outlet only) | 0 | 2501 | 1002 | 4101b | 5 | 5 | 0 | 0 | 0 | 0 | 0 | 4101b | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 2500 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 1 | 0.003125 | 0.003125 | 0 | 0 | 115200 | 0 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 47 | 4 | 1 | 200 | 2 | 5 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 20 | 0.2 | 0 | 0 | 0 | 0 | 0
ST-10.2 |  | Liq circ (ECS inlet to ICS outlet only) | On | On | Closed | Open | Open | Open | Open | Closed | Closed | Closed | Closed | Closed | Open | Closed | [5 empty] | 4102 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | Liq circ (ECS inlet to ICS outlet only) | 0 | 2501 | 1002 | 4102 | 5 | 5 | 0 | 0 | 0 | 0 | 0 | 4102 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 2500 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 1 | 0.003125 | 0.003125 | 0 | 0 | 115200 | 0 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 47 | 4 | 1 | 200 | 2 | 5 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 20 | 0.2 | 0 | 0 | 0 | 0 | 0
ST-10.3 |  | Liq circ (ECS inlet to ECS outlet only) | On | On | Closed | Closed | Open | Open | Open | Open | Closed | Closed | Closed | Closed | Open | Closed | [5 empty] | 4103 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | Liq circ (ECS inlet to ECS outlet only) | 0 | 2501 | 1002 | 4103 | 5 | 5 | 0 | 0 | 0 | 0 | 0 | 4103 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 2500 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 1 | 0.003125 | 0.003125 | 0 | 0 | 115200 | 0 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 47 | 4 | 1 | 200 | 2 | 5 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 20 | 0.2 | 0 | 0 | 0 | 0 | 0
ST-11.1 |  | Harvest loop fill | On | On | Closed | Closed | Closed | Open | Open | Open | Closed | Closed | Open | Closed | Open | Closed | [5 empty] | 4111 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | Harvest loop fill | 0 | 2501 | 1002 | 4111 | 100 | 100 | 0 | 0 | 0 | 0 | 0 | 4111 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 20000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 1 | 0.003125 | 0.003125 | 0 | 0 | 115200 | 0 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 47 | 4 | 1 | 200 | 2 | 5 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 20 | 0.2 | 0 | 0 | 0 | 0 | 0
ST-11.2 |  | Harvest loop in situ circulation | Off | On | Closed | Closed | Closed | Closed | Open | Open | Closed | Closed | Closed | Open | Natural | Natural | [5 empty] | 4112 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | Harvest loop in situ circulation | 0 | 2501 | 1002 | 4112 | 0 | 100 | 0 | 0 | 0 | 0 | 0 | 4112 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 60000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 1 | 0.003125 | 0.003125 | 0 | 0 | 115200 | 0 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 47 | 4 | 1 | 200 | 2 | 5 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 20 | 0.2 | 0 | 0 | 0 | 0 | 0
ST-11.3 |  | Harvest loop drain & cell removal | Off | On | Closed | Closed | Closed | Closed | Open | Open | Closed | Open | Closed | Open | Natural | Natural | [5 empty] | 4113 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | Harvest loop drain & cell removal | 0 | 2501 | 1002 | 4113 | 5 | 5 | 0 | 0 | 0 | 0 | 0 | 4113 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 2500 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 1 | 0.003125 | 0.003125 | 0 | 0 | 115200 | 0 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 47 | 4 | 1 | 200 | 2 | 5 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 20 | 0.2 | 0 | 0 | 0 | 0 | 0
ST-11.4 |  | Harvest loop flush | On | On | Closed | Closed | Closed | Open | Open | Open | Closed | Closed | Open | Closed | Natural | Natural | [5 empty] | 4114 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | Harvest loop flush | 0 | 2501 | 1002 | 4114 | 100 | 100 | 0 | 0 | 0 | 0 | 0 | 4114 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 60000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 1 | 0.003125 | 0.003125 | 0 | 0 | 115200 | 0 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 47 | 4 | 1 | 200 | 2 | 5 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 20 | 0.2 | 0 | 0 | 0 | 0 | 0
Current status of Cycler 1 as of 11/12/2023
[6 empty] | Valve | Plan | Actual | Tubing position
[6 empty] | V01 | NO | NO | Rear
[6 empty] | V02 | NO | NC | Front
[6 empty] | V03 | NO | W | Rear
[6 empty] | V04 | NC | W | Front
[6 empty] | V05 | NO | NO | Rear
[6 empty] | V06 | NC | W | Front
[6 empty] | V07 | NO | W | Rear
[6 empty] | V08 | NC | W | Front
[6 empty] | V09 | NC | W | Front
[6 empty] | V10 | NC | NC | Front
Note on CB2: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 [Tasks] There is a task anchored to this comment that cannot be viewed in your client. Comment: @Joe Nunn can you let me know the transition options that I can add here and the format/units you are working with for each transition type and I will start to think through what transition is appropriate for each state Reply: So currently I have Time, Temp and DO. I know those probably aren't that useful, but I can add additional transition functions relatively easily. Currently Time Transitions are in milliseconds, but I can change to seconds by just adding X1000 to the code, so whatever works best for you.
Note on AU4: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Set to -1 to used PID.
Note on Y5: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 [Tasks] There is a task anchored to this comment that cannot be viewed in your client. Comment: @Joe Nunn I have converted to the table to the left to the binary code required in the .csv config file and also assigned the associated ID in line with your nomenclature. Would you mind checking over this and then incorporating it into the config file? Reply: Looks great. The only change I would make is to not use a and b in stateIDs. if you need more digits, i can do that, but integers are much more efficient and easy to handle Reply: @Joe Nunn Yes I thought that might be an issue. The intention is to decide on either a or b once we have tested them, so this should disappear. Reply: @Scott Allan all good
Note on AU8: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Half the feed flow, as desire both retentate and permeate flow
Note on AU9: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Half the feed flow, as desire both retentate and permeate flow
Note on AU10: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Half the feed flow, as desire both retentate and permeate flow
Note on X19: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 [Tasks] There is a task anchored to this comment that cannot be viewed in your client. Comment: @Joe Nunn this will be the default cycling state

## Config
Pump | Valve
On | Open
Off | Closed
Rev | Natural

## Valve and pump states v2
States define the required "state change". i.e. if a state does not denote a valve as open or close then the previous state remains active.
State |  |  | P01 | P02 | V01 | V02 | V03 | V04 | V05 | V06 | V07 | V08 | V09 | V10 | V11 | V12 | [5 empty] | ID | P01 | P02 | V01 | V02 | V03 | V04 | V05 | V06 | V07 | V08 | V09 | V10 | V11 | V12 | [5 empty] | P1 (ml/min) | P2 (ml/min) |  | Actually NC - 0 & 1 swapped here |  | Actually W - does it need to be swapped? |  | Actually W - does it need to be swapped? | [6 empty] | Not used | Not used | Not used | Not used | Transition - Time
|  | Natural state (NO or NC) |  |  | NO | NO | NO | NC | NO | NC | NO | NC | NC | NC | NO | NC | [8 empty] | Rear | Front | Rear | Front | Rear | Front | Rear | Front | Front | Front | [9 empty] | NO | NC | W | W | NO | W | W | W | W | NC
|  | Actual natural state (NO or NC) |  |  | NO | NC | W | W | NO | W | W | W | W | NC | [10 empty] | NO | NC | W | W | NO | W | W | W | W | NC |  |  |  | NULLLEADER | MODID | COMMAND | STATEID | CIRCFLOW | PRESSUREFLOW | VALVE1 | VALVE2 | VALVE3 | VALVE4 | VALVE5 | VALVE6 | VALVE7 | VALVE8 | VALVE9 | VALVE10 | VALVE11 | VALVE12 | VALVE13 | VALVE14 | VALVE15 | VALVE16 | TRANSTIMEFLAG | TRANSTIMESP | PRESSURESP | OXYSP | OXYCIRCCHANNEL | CIRCPUMPCAL | PRESSUREPUMPCAL | PRESSUREKP | PRESSUREKI | PRESSUREKD | OXYKP | OXYKI | OXYKD | FLOWINSTALLED | PRESSUREINSTALLED | OXYINSTALLED | ACTIVEOXYCHANNEL1 | ACTIVEOXYCHANNEL2 | ACTIVEOXYCHANNEL3 | ACTIVEOXYCHANNEL4 | PUMP1CW | PUMP2CW | PUMP3CW | PUMP4CW | PUMP1DIRPIN | PUMP1STEPPIN | PUMP2DIRPIN | PUMP2STEPPIN | PUMP3DIRPIN | PUMP3STEPPIN | PUMP4DIRPIN | PUMP4STEPPIN | VALVE1PIN | VALVE2PIN | VALVE3PIN | VALVE4PIN | VALVE5PIN | VALVE6PIN | VALVE7PIN | VALVE8PIN | VALVE9PIN | VALVE10PIN | VALVE11PIN | VALVE12PIN | VALVE13PIN | VALVE14PIN | VALVE15PIN | VALVE16PIN | OXYCONFIGPARAMETER | PRESSURE0ADDRESS | PRESSURECHANNEL | FLOW0ADDRESS | DISPENSEPARA | DISPENSEVOLUMESP | PCBAUDRATE | OXYBAUDRATE | PRESSUREBAUDRATE | FLAGCHECK | NULLTRAILER
ST-01 |  | LEM reservoir fill | Off | Off | Natural | Natural | Natural | Natural | Natural | Natural | Natural | Natural | Natural | Natural | Natural | Natural | [5 empty] | 4010 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | LEM reservoir fill | 0 | 2501 | 1002 | 4010 | 5 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 2500 | 200 | 20 | 1 | 0.003125 | 0.003125 | 2 | 5 | 1 | 0.2 | 0 | 0 | 0 | 2 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 4 | 12 | 14 | 15 | 16 | 17 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4010 | 0
ST-02.1 |  | Line fill L04 and L11 | On | Off | Closed | Natural | Natural | Open | Closed | Natural | Closed | Open | Natural | Natural | Open | Closed | [5 empty] | 4021 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | Line fill L04 and L11 | 0 | 2501 | 1002 | 4021 | 40 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 10000 | 200 | 20 | 1 | 0.003125 | 0.003125 | 2 | 5 | 1 | 0.2 | 0 | 0 | 0 | 2 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 4 | 12 | 14 | 15 | 16 | 17 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4021 | 0
ST-02.2 |  | Line fill ECS | On | On | Closed | Closed | Closed | Open | Open | Open | Closed | Closed | Open | Closed | Open | Closed |  | Key | Description | NO | NC | 4022 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | Line fill ECS | 0 | 2501 | 1002 | 4022 | 50 | 50 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 60000 | 200 | 20 | 1 | 0.003125 | 0.003125 | 2 | 5 | 1 | 0.2 | 0 | 0 | 0 | 2 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 4 | 12 | 14 | 15 | 16 | 17 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4022 | 0
ST-02.3 |  | Line fill ICS | On | On | Open | Open | Closed | Closed | Open | Closed | Open | Closed | Closed | Open | Open | Closed |  | Natural | valve unpowered | 0 | 0 | 4023 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | Line fill ICS | 0 | 2501 | 1002 | 4023 | 100 | 100 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 20000 | 200 | 20 | 1 | 0.003125 | 0.003125 | 2 | 5 | 1 | 0.2 | 0 | 0 | 0 | 2 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 4 | 12 | 14 | 15 | 16 | 17 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4023 | 0
ST-02.4 |  | Line fill and circulation | On | On | Open | Open | Open | Closed | Open | Closed | Open | Closed | Closed | Closed | Open | Closed |  | Open | valve is open | 0 | 1 | 4024 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Line fill and circulation | 0 | 2501 | 1002 | 4024 | 100 | 100 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 40000 | 200 | 20 | 1 | 0.003125 | 0.003125 | 2 | 5 | 1 | 0.2 | 0 | 0 | 0 | 2 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 4 | 12 | 14 | 15 | 16 | 17 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4024 | 0
ST-03.1 |  | Drain circulation loop | On | On | Open | Open | Open | Closed | Open | Closed | Open | Closed | Closed | Closed | Closed | Open |  | Closed | valve is closed | 1 | 0 | 4031 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | Drain circulation loop | 0 | 2501 | 1002 | 4031 | 60 | 30 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 20000 | 200 | 20 | 1 | 0.003125 | 0.003125 | 2 | 5 | 1 | 0.2 | 0 | 0 | 0 | 2 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 4 | 12 | 14 | 15 | 16 | 17 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4031 | 0
ST-03.2 |  | Drain step 2 | On | On | Closed | Closed | Closed | Open | Open | Open | Closed | Closed | Open | Closed | Closed | Open |  | Rev | reversed |  |  | 4032 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | Drain step 2 | 0 | 2501 | 1002 | 4032 | 60 | 60 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 50000 | 200 | 20 | 1 | 0.003125 | 0.003125 | 2 | 5 | 1 | 0.2 | 0 | 0 | 0 | 2 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 4 | 12 | 14 | 15 | 16 | 17 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4032 | 0
ST-03.3 |  | Drain step 3 | On | Off | Closed | Natural | Natural | Open | Closed | Natural | Closed | Open | Natural | Closed | Closed | Open | [5 empty] | 4033 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | Drain step 3 | 0 | 2501 | 1002 | 4033 | 60 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 10000 | 200 | 20 | 1 | 0.003125 | 0.003125 | 2 | 5 | 1 | 0.2 | 0 | 0 | 0 | 2 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 4 | 12 | 14 | 15 | 16 | 17 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4033 | 0
ST-03.4 |  | Drain to empty the reservoir | On | Off | Closed | Natural | Closed | Open | Closed | Natural | Closed | Closed | Open | Open | Open | Closed | [5 empty] | 4034 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | Drain to empty the reservoir | 0 | 2501 | 1002 | 4034 | 60 | 0 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 20000 | 200 | 20 | 1 | 0.003125 | 0.003125 | 2 | 5 | 1 | 0.2 | 0 | 0 | 0 | 2 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 4 | 12 | 14 | 15 | 16 | 17 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4034 | 0
ST-03.5 |  | Drain ICS | On | On | Open | Open | Closed | Closed | Closed | Closed | Natural | Natural | Open | Closed | Open | Closed | [5 empty] | 4035 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Drain ICS | 0 | 2501 | 1002 | 4035 | 100 | 100 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 180000 | 200 | 20 | 1 | 0.003125 | 0.003125 | 2 | 5 | 1 | 0.2 | 0 | 0 | 0 | 2 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 4 | 12 | 14 | 15 | 16 | 17 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4035 | 0
ST-04 |  | Sampling | [19 empty] | 4040 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | Sampling | 0 | 2501 | 1002 | 4040 | 5 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 0 | 1 | 0 | 1 | 2500 | 200 | 20 | 1 | 0.003125 | 0.003125 | 2 | 5 | 1 | 0.2 | 0 | 0 | 0 | 2 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 4 | 12 | 14 | 15 | 16 | 17 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4040 | 0
ST-05 |  | Inoculation | Off | Off | Closed | Closed | Natural | Natural | Open | Open | Closed | Open | Natural | Closed | Natural | Natural | [5 empty] | 4050 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | Inoculation | 0 | 2501 | 1002 | 4050 | 5 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 2500 | 200 | 20 | 1 | 0.003125 | 0.003125 | 2 | 5 | 1 | 0.2 | 0 | 0 | 0 | 2 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 4 | 12 | 14 | 15 | 16 | 17 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4050 | 0
ST-06 |  | Dynamic seeding | Off | Off | Closed | Closed | Natural | Natural | Closed | Closed | Natural | Natural | Natural | Natural | Natural | Natural | [5 empty] | 4060 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Dynamic seeding | 0 | 2501 | 1002 | 4060 | 5 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 2500 | 200 | 20 | 1 | 0.003125 | 0.003125 | 2 | 5 | 1 | 0.2 | 0 | 0 | 0 | 2 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 4 | 12 | 14 | 15 | 16 | 17 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4060 | 0
ST-07 |  | ECS liquid drain/replacement (remove unattached cells) | On | On | Closed | Closed | Closed | Open | Open | Open | Natural | Natural | Open | Closed | Open | Closed | [5 empty] | 4070 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | ECS liquid drain/replacement (remove unattached cells) | 0 | 2501 | 1002 | 4070 | 5 | 5 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 2500 | 200 | 20 | 1 | 0.003125 | 0.003125 | 2 | 5 | 1 | 0.2 | 0 | 0 | 0 | 2 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 4 | 12 | 14 | 15 | 16 | 17 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4070 | 0
ST-08.1 |  | Liq circ (ICS inlet to ICS + ECS outlet only) | On | On | Open | Open | Open | Closed | Open | Closed | Open | Closed | Closed | Closed | Open | Closed | default state (most frequent) |  |  |  |  | 4081 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Liq circ (ICS inlet to ICS + ECS outlet only) | 0 | 2501 | 1002 | 4081 | 5 | 5 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 2500 | 200 | 20 | 1 | 0.003125 | 0.003125 | 2 | 5 | 1 | 0.2 | 0 | 0 | 0 | 2 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 4 | 12 | 14 | 15 | 16 | 17 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4081 | 0
ST-08.2 |  | Liq circ (ICS inlet to ICS outlet only) | On | On | Open | Open | Open | Closed | Closed | Closed | Natural | Natural | Closed | Closed | Open | Closed | [5 empty] | 4082 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Liq circ (ICS inlet to ICS outlet only) | 0 | 2501 | 1002 | 4082 | 60 | 60 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 10000 | 200 | 20 | 1 | 0.003125 | 0.003125 | 2 | 5 | 1 | 0.2 | 0 | 0 | 0 | 2 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 4 | 12 | 14 | 15 | 16 | 17 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4082 | 0
ST-08.3 |  | Liq circ (ICS inlet to ECS outlet only) | On | Off | Open | Closed | Natural | Closed | Open | Closed | Open | Closed | Natural | Closed | Open | Closed | [5 empty] | 4083 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Liq circ (ICS inlet to ECS outlet only) | 0 | 2501 | 1002 | 4083 | 5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 2500 | 200 | 20 | 1 | 0.003125 | 0.003125 | 2 | 5 | 1 | 0.2 | 0 | 0 | 0 | 2 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 4 | 12 | 14 | 15 | 16 | 17 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4083 | 0
ST-09.1a | a | Liq circ (ICS + ECS inlet to ICS + ECS outlet) - option 1 | Rev | Rev | Open | Open | Open | Closed | Open | Open | Open | Closed | Closed | Closed | Open | Closed | [5 empty] | 4091a | -1 | -1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | Liq circ (ICS + ECS inlet to ICS + ECS outlet) - option 1 | 0 | 2501 | 1002 | 4091a | 5 | 5 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 2500 | 200 | 20 | 1 | 0.003125 | 0.003125 | 2 | 5 | 1 | 0.2 | 0 | 0 | 0 | 2 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 4 | 12 | 14 | 15 | 16 | 17 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4091a | 0
ST-09.1b | b | Liq circ (ICS + ECS inlet to ICS + ECS outlet) - option 2 | On | On | Open | Open | Open | Open | Open | Open | Closed | Closed | Closed | Closed | Open | Closed | [5 empty] | 4091b | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | Liq circ (ICS + ECS inlet to ICS + ECS outlet) - option 2 | 0 | 2501 | 1002 | 4091b | 5 | 5 | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 2500 | 200 | 20 | 1 | 0.003125 | 0.003125 | 2 | 5 | 1 | 0.2 | 0 | 0 | 0 | 2 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 4 | 12 | 14 | 15 | 16 | 17 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4091b | 0
ST-09.2 |  | Liq circ (ICS + ECS inlet to ICS outlet only) | On | On | Open | Open | Open | Open | Open | Closed | Closed | Closed | Closed | Closed | Open | Closed | [5 empty] | 4092 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | Liq circ (ICS + ECS inlet to ICS outlet only) | 0 | 2501 | 1002 | 4092 | 5 | 5 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 2500 | 200 | 20 | 1 | 0.003125 | 0.003125 | 2 | 5 | 1 | 0.2 | 0 | 0 | 0 | 2 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 4 | 12 | 14 | 15 | 16 | 17 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4092 | 0
ST-09.3 |  | Liq circ (ICS + ECS inlet to ECS outlet only) | On | On | Open | Closed | Open | Open | Open | Open | Closed | Closed | Closed | Closed | Open | Closed | [5 empty] | 4093 | 1 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | Liq circ (ICS + ECS inlet to ECS outlet only) | 0 | 2501 | 1002 | 4093 | 5 | 5 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 2500 | 200 | 20 | 1 | 0.003125 | 0.003125 | 2 | 5 | 1 | 0.2 | 0 | 0 | 0 | 2 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 4 | 12 | 14 | 15 | 16 | 17 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4093 | 0
ST-10.1a | a | Liq circ (ECS inlet to ICS + ECS outlet only) | Rev | Rev | Open | Closed | Open | Closed | Open | Open | Open | Closed | Closed | Closed | Open | Closed | [5 empty] | 4101a | -1 | -1 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | Liq circ (ECS inlet to ICS + ECS outlet only) | 0 | 2501 | 1002 | 4101a | 5 | 5 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 2500 | 200 | 20 | 1 | 0.003125 | 0.003125 | 2 | 5 | 1 | 0.2 | 0 | 0 | 0 | 2 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 4 | 12 | 14 | 15 | 16 | 17 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4101a | 0
ST-10.1b | b | Liq circ (ECS inlet to ICS + ECS outlet only) | On | On | Closed | Open | Open | Open | Open | Open | Closed | Closed | Closed | Closed | Open | Closed | [5 empty] | 4101b | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | Liq circ (ECS inlet to ICS + ECS outlet only) | 0 | 2501 | 1002 | 4101b | 5 | 5 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 2500 | 200 | 20 | 1 | 0.003125 | 0.003125 | 2 | 5 | 1 | 0.2 | 0 | 0 | 0 | 2 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 4 | 12 | 14 | 15 | 16 | 17 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4101b | 0
ST-10.2 |  | Liq circ (ECS inlet to ICS outlet only) | On | On | Closed | Open | Open | Open | Open | Closed | Closed | Closed | Closed | Closed | Open | Closed | [5 empty] | 4102 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | Liq circ (ECS inlet to ICS outlet only) | 0 | 2501 | 1002 | 4102 | 5 | 5 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 2500 | 200 | 20 | 1 | 0.003125 | 0.003125 | 2 | 5 | 1 | 0.2 | 0 | 0 | 0 | 2 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 4 | 12 | 14 | 15 | 16 | 17 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4102 | 0
ST-10.3 |  | Liq circ (ECS inlet to ECS outlet only) | On | On | Closed | Closed | Open | Open | Open | Open | Closed | Closed | Closed | Closed | Open | Closed | [5 empty] | 4103 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | Liq circ (ECS inlet to ECS outlet only) | 0 | 2501 | 1002 | 4103 | 5 | 5 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 2500 | 200 | 20 | 1 | 0.003125 | 0.003125 | 2 | 5 | 1 | 0.2 | 0 | 0 | 0 | 2 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 4 | 12 | 14 | 15 | 16 | 17 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4103 | 0
ST-11.1 |  | Harvest loop fill | On | On | Closed | Closed | Closed | Open | Open | Open | Closed | Closed | Open | Closed | Open | Closed | [5 empty] | 4111 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | Harvest loop fill | 0 | 2501 | 1002 | 4111 | 100 | 100 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 20000 | 200 | 20 | 1 | 0.003125 | 0.003125 | 2 | 5 | 1 | 0.2 | 0 | 0 | 0 | 2 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 4 | 12 | 14 | 15 | 16 | 17 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4111 | 0
ST-11.2 |  | Harvest loop in situ circulation | Off | On | Closed | Closed | Closed | Closed | Open | Open | Closed | Closed | Closed | Open | Natural | Natural | [5 empty] | 4112 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | Harvest loop in situ circulation | 0 | 2501 | 1002 | 4112 | 0 | 100 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 60000 | 200 | 20 | 1 | 0.003125 | 0.003125 | 2 | 5 | 1 | 0.2 | 0 | 0 | 0 | 2 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 4 | 12 | 14 | 15 | 16 | 17 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4112 | 0
ST-11.3 |  | Harvest loop drain & cell removal | Off | On | Closed | Closed | Closed | Closed | Open | Open | Closed | Open | Closed | Open | Natural | Natural | [5 empty] | 4113 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | Harvest loop drain & cell removal | 0 | 2501 | 1002 | 4113 | 5 | 5 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 2500 | 200 | 20 | 1 | 0.003125 | 0.003125 | 2 | 5 | 1 | 0.2 | 0 | 0 | 0 | 2 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 4 | 12 | 14 | 15 | 16 | 17 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4113 | 0
ST-11.4 |  | Harvest loop flush | On | On | Closed | Closed | Closed | Open | Open | Open | Closed | Closed | Open | Closed | Natural | Natural | [5 empty] | 4114 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | Harvest loop flush | 0 | 2501 | 1002 | 4114 | 100 | 100 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 60000 | 200 | 20 | 1 | 0.003125 | 0.003125 | 2 | 5 | 1 | 0.2 | 0 | 0 | 0 | 2 | 1 | 1 | 0 | 0 | 0 |  |  |  |  | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 13 | 2 | 3 | 11 | 5 | 6 | 7 | 8 | 9 | 10 | 4 | 12 | 14 | 15 | 16 | 17 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4114 | 0
Current status of Cycler 1 as of 11/12/2023
[6 empty] | Valve | Plan | Actual | Tubing position
[6 empty] | V01 | NO | NO | Rear
[6 empty] | V02 | NO | NC | Front
[6 empty] | V03 | NO | W | Rear
[6 empty] | V04 | NC | W | Front
[6 empty] | V05 | NO | NO | Rear
[6 empty] | V06 | NC | W | Front
[6 empty] | V07 | NO | W | Rear
[6 empty] | V08 | NC | W | Front
[6 empty] | V09 | NC | W | Front
[6 empty] | V10 | NC | NC | Front
Note on BL2: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 [Tasks] There is a task anchored to this comment that cannot be viewed in your client. Comment: @Joe Nunn can you let me know the transition options that I can add here and the format/units you are working with for each transition type and I will start to think through what transition is appropriate for each state Reply: So currently I have Time, Temp and DO. I know those probably aren't that useful, but I can add additional transition functions relatively easily. Currently Time Transitions are in milliseconds, but I can change to seconds by just adding X1000 to the code, so whatever works best for you.
Note on AU4: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Set to -1 to used PID.
Note on Y5: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 [Tasks] There is a task anchored to this comment that cannot be viewed in your client. Comment: @Joe Nunn I have converted to the table to the left to the binary code required in the .csv config file and also assigned the associated ID in line with your nomenclature. Would you mind checking over this and then incorporating it into the config file? Reply: Looks great. The only change I would make is to not use a and b in stateIDs. if you need more digits, i can do that, but integers are much more efficient and easy to handle Reply: @Joe Nunn Yes I thought that might be an issue. The intention is to decide on either a or b once we have tested them, so this should disappear. Reply: @Scott Allan all good
Note on AU8: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Half the feed flow, as desire both retentate and permeate flow
Note on AU9: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Half the feed flow, as desire both retentate and permeate flow
Note on AU10: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Half the feed flow, as desire both retentate and permeate flow
Note on X19: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 [Tasks] There is a task anchored to this comment that cannot be viewed in your client. Comment: @Joe Nunn this will be the default cycling state
Valid until 28/01/2024. Then transition to v3.

## Valve and pump states v3
States define the required "state change". i.e. if a state does not denote a valve as open or close then the previous state remains active. | [41 empty] | P1 (ml/min) | P2 (ml/min) |  | Actually NC - 0 & 1 swapped here |  | Actually W - does it need to be swapped? |  | Actually W - does it need to be swapped? | [6 empty] | Not used | Not used | Not used | Not used | Transition - Time
State |  |  | P01 | P02 | V01 | V02 | V03 | V04 | V05 | V06 | V07 | V08 | V09 | V10 | V11 | V12 | [5 empty] | ID | P01 | P02 | V01 | V02 | V03 | V04 | V05 | V06 | V07 | V08 | V09 | V10 | V11 | V12 | [7 empty] | NO | NC | W | W | NO | W | W | W | W | NC
|  | Natural state (NO or NC) |  |  | NO | NO | NO | NC | NO | NC | NO | NC | NC | NC | NO | NC | [8 empty] | Rear | Front | Rear | Front | Rear | Front | Rear | Front | Front | Front |  |  |  | nullLeader | modId | command | stateId | circFlow | pressureFlow | valve1 | valve2 | valve3 | valve4 | valve5 | valve6 | valve7 | valve8 | valve9 | valve10 | valve11 | valve12 | valve13 | valve14 | valve15 | valve16 | transTimeFlag | transTimeSp | pressureSP | oxySP | presInterval | flowInterval | oxyInterval | reportInterval | oxyTemp | oxyCircChannel | pumpSpeedRatio | circPumpCal | pressurePumpCal | pressureKp | pressureKi | pressureKd | oxyKp | oxyKi | oxyKd | flowInstalled | pressureInstalled | oxyInstalled | activeOxyChannel1 | activeOxyChannel2 | activeOxyChannel3 | activeOxyChannel4 | pump1Cw | pump2Cw | pump3Cw | pump4Cw | pump1DirPin | pump1StepPin | pump2DirPin | pump2StepPin | pump3DirPin | pump3StepPin | pump4DirPin | pump4StepPin | valve1Pin | valve2Pin | valve3Pin | valve4Pin | valve5Pin | valve6Pin | valve7Pin | valve8Pin | valve9Pin | valve10Pin | valve11Pin | valve12Pin | valve13Pin | valve14Pin | valve15Pin | valve16Pin | oxyConfigParameter | pressure0Address | pressureChannel | flow0Address | dispensePara | dispenseVolumeSP | pcBaudRate | oxyBaudRate | pressureBaudRate | flagCheck | nullTrailer
|  | Actual natural state (NO or NC) |  |  | NO | NC | W | W | NO | W | W | W | W | NC | [10 empty] | NO | NC | W | W | NO | W | W | W | W | NC |  |  |  | int | int | int | int | double | double | bool | bool | bool | bool | bool | bool | bool | bool | bool | bool | bool | bool | bool | bool | bool | bool | bool | long | double | double | double | double | double | double | float | int | float | float | float | double | double | double | double | double | double | int | int | int | int | int | int | int | bool | bool | bool | bool | int | int | int | int | int | int | int | int | int | int | int | int | int | int | int | int | int | int | int | int | int | int | int | int | int | int | int | int | int | int | double | double | double | int | int
ST-01 |  | LEM reservoir fill | Off | Off | Natural | Natural | Natural | Natural | Natural | Natural | Natural | Natural | Natural | Natural | Natural | Natural | [5 empty] | 4010 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | LEM reservoir fill | 0 | 2501 | 1002 | 4010 | 5 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 2500 | 200 | 20 | 100 | 100 | 60000 | 1000 | 37 | 1 | 0.9 | 0.003125 | 0.003125 | 2 | 5 | 1 | 1 | 2 | 0 | 1 | 2 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 11 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 24 | 26 | 28 | 30 | 32 | 22 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4010 | 0
ST-02.1 |  | Line fill L04 and L11 | On | Off | Closed | Natural | Natural | Open | Closed | Natural | Closed | Open | Natural | Natural | Open | Closed | [5 empty] | 4021 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | Line fill L04 and L11 | 0 | 2501 | 1002 | 4021 | 40 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 15000 | 200 | 20 | 100 | 100 | 60000 | 1000 | 37 | 1 | 0.9 | 0.003125 | 0.003125 | 2 | 5 | 1 | 1 | 2 | 0 | 1 | 2 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 11 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 24 | 26 | 28 | 30 | 32 | 22 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4021 | 0
ST-02.2 |  | Line fill ECS | On | On | Closed | Closed | Closed | Open | Open | Open | Closed | Closed | Open | Closed | Open | Closed |  | Key | Description | NO | NC | 4022 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | Line fill ECS | 0 | 2501 | 1002 | 4022 | 50 | 50 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 60000 | 200 | 20 | 100 | 100 | 60000 | 1000 | 37 | 1 | 0.9 | 0.003125 | 0.003125 | 2 | 5 | 1 | 1 | 2 | 0 | 1 | 2 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 11 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 24 | 26 | 28 | 30 | 32 | 22 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4022 | 0
ST-02.3 |  | Line fill ICS | On | On | Open | Open | Closed | Closed | Open | Closed | Open | Closed | Closed | Open | Open | Closed |  | Natural | valve unpowered | 0 | 0 | 4023 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | Line fill ICS | 0 | 2501 | 1002 | 4023 | 100 | 100 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 20000 | 200 | 20 | 100 | 100 | 60000 | 1000 | 37 | 1 | 0.9 | 0.003125 | 0.003125 | 2 | 5 | 1 | 1 | 2 | 0 | 1 | 2 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 11 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 24 | 26 | 28 | 30 | 32 | 22 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4023 | 0
ST-02.4 |  | Line fill and circulation | On | On | Open | Open | Open | Closed | Open | Closed | Open | Closed | Closed | Closed | Open | Closed |  | Open | valve is open | 0 | 1 | 4024 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Line fill and circulation | 0 | 2501 | 1002 | 4024 | 100 | 100 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 40000 | 200 | 20 | 100 | 100 | 60000 | 1000 | 37 | 1 | 0.9 | 0.003125 | 0.003125 | 2 | 5 | 1 | 1 | 2 | 0 | 1 | 2 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 11 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 24 | 26 | 28 | 30 | 32 | 22 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4024 | 0
ST-03.1 |  | Drain circulation loop | On | On | Open | Open | Open | Closed | Open | Closed | Open | Closed | Closed | Closed | Closed | Open |  | Closed | valve is closed | 1 | 0 | 4031 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | Drain circulation loop | 0 | 2501 | 1002 | 4031 | 60 | 30 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 20000 | 200 | 20 | 100 | 100 | 60000 | 1000 | 37 | 1 | 0.9 | 0.003125 | 0.003125 | 2 | 5 | 1 | 1 | 2 | 0 | 1 | 2 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 11 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 24 | 26 | 28 | 30 | 32 | 22 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4031 | 0
ST-03.2 |  | Drain step 2 | On | On | Closed | Closed | Closed | Open | Open | Open | Closed | Closed | Open | Closed | Closed | Open |  | Rev | reversed |  |  | 4032 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | Drain step 2 | 0 | 2501 | 1002 | 4032 | 60 | 60 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 60000 | 200 | 20 | 100 | 100 | 60000 | 1000 | 37 | 1 | 0.9 | 0.003125 | 0.003125 | 2 | 5 | 1 | 1 | 2 | 0 | 1 | 2 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 11 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 24 | 26 | 28 | 30 | 32 | 22 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4032 | 0
ST-03.3 |  | Drain step 3 | On | Off | Closed | Natural | Natural | Open | Closed | Natural | Closed | Open | Natural | Closed | Closed | Open | [5 empty] | 4033 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | Drain step 3 | 0 | 2501 | 1002 | 4033 | 60 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 10000 | 200 | 20 | 100 | 100 | 60000 | 1000 | 37 | 1 | 0.9 | 0.003125 | 0.003125 | 2 | 5 | 1 | 1 | 2 | 0 | 1 | 2 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 11 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 24 | 26 | 28 | 30 | 32 | 22 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4033 | 0
ST-03.4 |  | Drain to empty the reservoir | On | Off | Closed | Natural | Closed | Open | Closed | Natural | Closed | Closed | Open | Open | Open | Closed | [5 empty] | 4034 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | Drain to empty the reservoir | 0 | 2501 | 1002 | 4034 | 60 | 0 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 20000 | 200 | 20 | 100 | 100 | 60000 | 1000 | 37 | 1 | 0.9 | 0.003125 | 0.003125 | 2 | 5 | 1 | 1 | 2 | 0 | 1 | 2 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 11 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 24 | 26 | 28 | 30 | 32 | 22 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4034 | 0
ST-03.5 |  | Drain ICS | On | On | Open | Open | Closed | Closed | Closed | Closed | Natural | Natural | Open | Closed | Open | Closed | [5 empty] | 4035 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Drain ICS | 0 | 2501 | 1002 | 4035 | 100 | 100 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 180000 | 200 | 20 | 100 | 100 | 60000 | 1000 | 37 | 1 | 0.9 | 0.003125 | 0.003125 | 2 | 5 | 1 | 1 | 2 | 0 | 1 | 2 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 11 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 24 | 26 | 28 | 30 | 32 | 22 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4035 | 0
ST-04 |  | Sampling | [19 empty] | 4040 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | Sampling | 0 | 2501 | 1002 | 4040 | 5 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 0 | 1 | 0 | 1 | 2500 | 200 | 20 | 100 | 100 | 60000 | 1000 | 37 | 1 | 0.9 | 0.003125 | 0.003125 | 2 | 5 | 1 | 1 | 2 | 0 | 1 | 2 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 11 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 24 | 26 | 28 | 30 | 32 | 22 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4040 | 0
ST-05 |  | Inoculation | Off | Off | Closed | Closed | Natural | Natural | Open | Open | Closed | Open | Natural | Closed | Natural | Natural | [5 empty] | 4050 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | Inoculation | 0 | 2501 | 1002 | 4050 | 5 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 2500 | 200 | 20 | 100 | 100 | 60000 | 1000 | 37 | 1 | 0.9 | 0.003125 | 0.003125 | 2 | 5 | 1 | 1 | 2 | 0 | 1 | 2 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 11 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 24 | 26 | 28 | 30 | 32 | 22 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4050 | 0
ST-06 |  | Dynamic seeding | Off | Off | Closed | Closed | Natural | Natural | Closed | Closed | Natural | Natural | Natural | Natural | Natural | Natural | [5 empty] | 4060 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Dynamic seeding | 0 | 2501 | 1002 | 4060 | 5 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 2500 | 200 | 20 | 100 | 100 | 60000 | 1000 | 37 | 1 | 0.9 | 0.003125 | 0.003125 | 2 | 5 | 1 | 1 | 2 | 0 | 1 | 2 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 11 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 24 | 26 | 28 | 30 | 32 | 22 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4060 | 0
ST-07 |  | ECS liquid drain/replacement (remove unattached cells) | On | On | Closed | Closed | Closed | Open | Open | Open | Natural | Natural | Open | Closed | Open | Closed | [5 empty] | 4070 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | ECS liquid drain/replacement (remove unattached cells) | 0 | 2501 | 1002 | 4070 | 5 | 5 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 2500 | 200 | 20 | 100 | 100 | 60000 | 1000 | 37 | 1 | 0.9 | 0.003125 | 0.003125 | 2 | 5 | 1 | 1 | 2 | 0 | 1 | 2 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 11 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 24 | 26 | 28 | 30 | 32 | 22 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4070 | 0
ST-08.1 |  | Liq circ (ICS inlet to ICS + ECS outlet only) | On | On | Open | Open | Open | Closed | Open | Closed | Open | Closed | Closed | Closed | Open | Closed | default state (most frequent) |  |  |  |  | 4081 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Liq circ (ICS inlet to ICS + ECS outlet only) | 0 | 2501 | 1002 | 4081 | 5 | 5 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 2500 | 200 | 20 | 100 | 100 | 60000 | 1000 | 37 | 1 | 0.9 | 0.003125 | 0.003125 | 2 | 5 | 1 | 1 | 2 | 0 | 1 | 2 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 11 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 24 | 26 | 28 | 30 | 32 | 22 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4081 | 0
ST-08.2 |  | Liq circ (ICS inlet to ICS outlet only) | On | On | Open | Open | Open | Closed | Closed | Closed | Natural | Natural | Closed | Closed | Open | Closed | [5 empty] | 4082 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Liq circ (ICS inlet to ICS outlet only) | 0 | 2501 | 1002 | 4082 | 60 | 60 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 10000 | 200 | 20 | 100 | 100 | 60000 | 1000 | 37 | 1 | 0.9 | 0.003125 | 0.003125 | 2 | 5 | 1 | 1 | 2 | 0 | 1 | 2 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 11 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 24 | 26 | 28 | 30 | 32 | 22 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4082 | 0
ST-08.3 |  | Liq circ (ICS inlet to ECS outlet only) | On | Off | Open | Closed | Natural | Closed | Open | Closed | Open | Closed | Natural | Closed | Open | Closed | [5 empty] | 4083 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Liq circ (ICS inlet to ECS outlet only) | 0 | 2501 | 1002 | 4083 | 5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 2500 | 200 | 20 | 100 | 100 | 60000 | 1000 | 37 | 1 | 0.9 | 0.003125 | 0.003125 | 2 | 5 | 1 | 1 | 2 | 0 | 1 | 2 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 11 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 24 | 26 | 28 | 30 | 32 | 22 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4083 | 0
ST-09.1a | a | Liq circ (ICS + ECS inlet to ICS + ECS outlet) - option 1 | Rev | Rev | Open | Open | Open | Closed | Open | Open | Open | Closed | Closed | Closed | Open | Closed | [5 empty] | 4091a | -1 | -1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | Liq circ (ICS + ECS inlet to ICS + ECS outlet) - option 1 | 0 | 2501 | 1002 | 4091a | 5 | 5 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 2500 | 200 | 20 | 100 | 100 | 60000 | 1000 | 37 | 1 | 0.9 | 0.003125 | 0.003125 | 2 | 5 | 1 | 1 | 2 | 0 | 1 | 2 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 11 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 24 | 26 | 28 | 30 | 32 | 22 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4091a | 0
ST-09.1b | b | Liq circ (ICS + ECS inlet to ICS + ECS outlet) - option 2 | On | On | Open | Open | Open | Open | Open | Open | Closed | Closed | Closed | Closed | Open | Closed | [5 empty] | 4091b | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | Liq circ (ICS + ECS inlet to ICS + ECS outlet) - option 2 | 0 | 2501 | 1002 | 4091b | 5 | 5 | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 2500 | 200 | 20 | 100 | 100 | 60000 | 1000 | 37 | 1 | 0.9 | 0.003125 | 0.003125 | 2 | 5 | 1 | 1 | 2 | 0 | 1 | 2 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 11 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 24 | 26 | 28 | 30 | 32 | 22 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4091b | 0
ST-09.2 |  | Liq circ (ICS + ECS inlet to ICS outlet only) | On | On | Open | Open | Open | Open | Open | Closed | Closed | Closed | Closed | Closed | Open | Closed | [5 empty] | 4092 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | Liq circ (ICS + ECS inlet to ICS outlet only) | 0 | 2501 | 1002 | 4092 | 5 | 5 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 2500 | 200 | 20 | 100 | 100 | 60000 | 1000 | 37 | 1 | 0.9 | 0.003125 | 0.003125 | 2 | 5 | 1 | 1 | 2 | 0 | 1 | 2 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 11 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 24 | 26 | 28 | 30 | 32 | 22 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4092 | 0
ST-09.3 |  | Liq circ (ICS + ECS inlet to ECS outlet only) | On | On | Open | Closed | Open | Open | Open | Open | Closed | Closed | Closed | Closed | Open | Closed | [5 empty] | 4093 | 1 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | Liq circ (ICS + ECS inlet to ECS outlet only) | 0 | 2501 | 1002 | 4093 | 5 | 5 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 2500 | 200 | 20 | 100 | 100 | 60000 | 1000 | 37 | 1 | 0.9 | 0.003125 | 0.003125 | 2 | 5 | 1 | 1 | 2 | 0 | 1 | 2 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 11 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 24 | 26 | 28 | 30 | 32 | 22 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4093 | 0
ST-10.1a | a | Liq circ (ECS inlet to ICS + ECS outlet only) | Rev | Rev | Open | Closed | Open | Closed | Open | Open | Open | Closed | Closed | Closed | Open | Closed | [5 empty] | 4101a | -1 | -1 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | Liq circ (ECS inlet to ICS + ECS outlet only) | 0 | 2501 | 1002 | 4101a | 5 | 5 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 2500 | 200 | 20 | 100 | 100 | 60000 | 1000 | 37 | 1 | 0.9 | 0.003125 | 0.003125 | 2 | 5 | 1 | 1 | 2 | 0 | 1 | 2 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 11 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 24 | 26 | 28 | 30 | 32 | 22 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4101a | 0
ST-10.1b | b | Liq circ (ECS inlet to ICS + ECS outlet only) | On | On | Closed | Open | Open | Open | Open | Open | Closed | Closed | Closed | Closed | Open | Closed | [5 empty] | 4101b | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | Liq circ (ECS inlet to ICS + ECS outlet only) | 0 | 2501 | 1002 | 4101b | 5 | 5 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 2500 | 200 | 20 | 100 | 100 | 60000 | 1000 | 37 | 1 | 0.9 | 0.003125 | 0.003125 | 2 | 5 | 1 | 1 | 2 | 0 | 1 | 2 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 11 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 24 | 26 | 28 | 30 | 32 | 22 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4101b | 0
ST-10.2 |  | Liq circ (ECS inlet to ICS outlet only) | On | On | Closed | Open | Open | Open | Open | Closed | Closed | Closed | Closed | Closed | Open | Closed | [5 empty] | 4102 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | Liq circ (ECS inlet to ICS outlet only) | 0 | 2501 | 1002 | 4102 | 5 | 5 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 2500 | 200 | 20 | 100 | 100 | 60000 | 1000 | 37 | 1 | 0.9 | 0.003125 | 0.003125 | 2 | 5 | 1 | 1 | 2 | 0 | 1 | 2 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 11 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 24 | 26 | 28 | 30 | 32 | 22 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4102 | 0
ST-10.3 |  | Liq circ (ECS inlet to ECS outlet only) | On | On | Closed | Closed | Open | Open | Open | Open | Closed | Closed | Closed | Closed | Open | Closed | [5 empty] | 4103 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | Liq circ (ECS inlet to ECS outlet only) | 0 | 2501 | 1002 | 4103 | 5 | 5 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 2500 | 200 | 20 | 100 | 100 | 60000 | 1000 | 37 | 1 | 0.9 | 0.003125 | 0.003125 | 2 | 5 | 1 | 1 | 2 | 0 | 1 | 2 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 11 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 24 | 26 | 28 | 30 | 32 | 22 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4103 | 0
ST-11.1 |  | Harvest loop fill | On | On | Closed | Closed | Closed | Open | Open | Open | Closed | Closed | Open | Closed | Open | Closed | [5 empty] | 4111 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | Harvest loop fill | 0 | 2501 | 1002 | 4111 | 100 | 100 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 20000 | 200 | 20 | 100 | 100 | 60000 | 1000 | 37 | 1 | 0.9 | 0.003125 | 0.003125 | 2 | 5 | 1 | 1 | 2 | 0 | 1 | 2 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 11 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 24 | 26 | 28 | 30 | 32 | 22 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4111 | 0
ST-11.2 |  | Harvest loop in situ circulation | Off | On | Closed | Closed | Closed | Closed | Open | Open | Closed | Closed | Closed | Open | Natural | Natural | [5 empty] | 4112 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | Harvest loop in situ circulation | 0 | 2501 | 1002 | 4112 | 0 | 100 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 60000 | 200 | 20 | 100 | 100 | 60000 | 1000 | 37 | 1 | 0.9 | 0.003125 | 0.003125 | 2 | 5 | 1 | 1 | 2 | 0 | 1 | 2 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 11 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 24 | 26 | 28 | 30 | 32 | 22 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4112 | 0
ST-11.3 |  | Harvest loop drain & cell removal | Off | On | Closed | Closed | Closed | Closed | Open | Open | Closed | Open | Closed | Open | Natural | Natural | [5 empty] | 4113 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | Harvest loop drain & cell removal | 0 | 2501 | 1002 | 4113 | 5 | 5 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 2500 | 200 | 20 | 100 | 100 | 60000 | 1000 | 37 | 1 | 0.9 | 0.003125 | 0.003125 | 2 | 5 | 1 | 1 | 2 | 0 | 1 | 2 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 11 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 24 | 26 | 28 | 30 | 32 | 22 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4113 | 0
ST-11.4 |  | Harvest loop flush | On | On | Closed | Closed | Closed | Open | Open | Open | Closed | Closed | Open | Closed | Natural | Natural | [5 empty] | 4114 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | Harvest loop flush | 0 | 2501 | 1002 | 4114 | 100 | 100 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 60000 | 200 | 20 | 100 | 100 | 60000 | 1000 | 37 | 1 | 0.9 | 0.003125 | 0.003125 | 2 | 5 | 1 | 1 | 2 | 0 | 1 | 2 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 27 | 29 | 31 | 33 | 39 | 40 | 41 | 42 | 11 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 24 | 26 | 28 | 30 | 32 | 22 | 47 | 54 | 0 | 8 | 0 | 0 | 115200 | 115200 | 115200 | 4114 | 0
Current status of Cycler 1 as of 11/12/2023
[6 empty] | Valve | Plan | Actual | Tubing position
[6 empty] | V01 | NO | NO | Rear
[6 empty] | V02 | NO | NC | Front
[6 empty] | V03 | NO | W | Rear
[6 empty] | V04 | NC | W | Front
[6 empty] | V05 | NO | NO | Rear
[6 empty] | V06 | NC | W | Front
[6 empty] | V07 | NO | W | Rear
[6 empty] | V08 | NC | W | Front
[6 empty] | V09 | NC | W | Front
[6 empty] | V10 | NC | NC | Front
Note on BL1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 [Tasks] There is a task anchored to this comment that cannot be viewed in your client. Comment: @Joe Nunn can you let me know the transition options that I can add here and the format/units you are working with for each transition type and I will start to think through what transition is appropriate for each state Reply: So currently I have Time, Temp and DO. I know those probably aren't that useful, but I can add additional transition functions relatively easily. Currently Time Transitions are in milliseconds, but I can change to seconds by just adding X1000 to the code, so whatever works best for you.
Note on AU3: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Set to -1 to used PID.
Note on Y5: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 [Tasks] There is a task anchored to this comment that cannot be viewed in your client. Comment: @Joe Nunn I have converted to the table to the left to the binary code required in the .csv config file and also assigned the associated ID in line with your nomenclature. Would you mind checking over this and then incorporating it into the config file? Reply: Looks great. The only change I would make is to not use a and b in stateIDs. if you need more digits, i can do that, but integers are much more efficient and easy to handle Reply: @Joe Nunn Yes I thought that might be an issue. The intention is to decide on either a or b once we have tested them, so this should disappear. Reply: @Scott Allan all good
Note on AU8: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Half the feed flow, as desire both retentate and permeate flow
Note on AU9: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Half the feed flow, as desire both retentate and permeate flow
Note on AU10: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Half the feed flow, as desire both retentate and permeate flow
Note on X19: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 [Tasks] There is a task anchored to this comment that cannot be viewed in your client. Comment: @Joe Nunn this will be the default cycling state

## Not used in v2
Transition - Temp |  | Transition - DO
TRANSTEMPFLAG | TRANSTEMPSP | TRANSOXYFLAG | TRANSOXYSP | SERIALERRORFLAG | TRANSTYPE | NA | NA | NA
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0